## Slide 1: Environmental Magnetic Fields for the Neutron Electric Dipole Moment Experiment at TRIUMF
By: Maedeh Lavvaf
Magnetically Shielded Room (MSR)
Mapping the magnetic field in TUCAN area

## Slide 2: nEDM Experiment Motivation
Helps explain matter-antimatter asymmetry
Under the time reversal, is reversed, but is not.
T violation CP violation under CPT symmetry
Beyond Standard Model
Standard Model
Current limit
TUCAN target by a factor of 30
Notes: SM doesn’t contain enough CP violation in order to explain particle creation. The non-zero nEDM is a source of CP violation that could explain baryon asymmetry. If the nEDM found, it could contribute to our understanding of extensions of the SM. The TRIUMF colaboration aims to improve the sensitivity by a factor of 30 with a goal of sensitivity of 10^-27e.cm

## Slide 3: TUCAN nEDM Experiment
Goal: 10^-27 e.cm within 400 measurement days
Experiment | Timeline
Detailed design and build | Now -2022
Installation and Commissioning nEDM apparatus nEDM run | 2023

## Slide 4: TUCAN Magnetically Shielded Room
A Magnetically Shielded Room (MSR) is a multi-layer room built with high permeability material, such as mumetal, and high conductivity sheet materials, such as copper.
Side View

## Slide 5: MSR Outer Layer Saturation
Need to use actual field at TRIUMF site
Unique Challenge: Saturation (0.7 T) of the MSR most outer layer from the TRIUMF Cyclotron fringe field
Opera FEA (static mu=1M)
Using B=350 µT Vertical Field
MSR outerlayer > 0.1 T, especially near the holes
Saturation near big holes
0.95
0.8
0.7
0.6
0.5
0.4
0.3
0.2
0.1
0.18E-3
Surface B field (T)
Surface B field (T)
1.3E-6
0.32
0.25
0.20
0.15
0.10
0.05
Opera FEA (static mu=1M)
Using B=350 µT Vertical Field

## Slide 6: Field Mapping with TRIUMF cyclotron on/off
MSR foot print
Mapping foot print
Map area:
Xtot =400cm
Ytot =500cm
Ztot =350cm
in 40 cm increments

## Slide 7: Field Mapping with TRIUMF cyclotron on/off
120 µT
505 µT
Cyclotron Off
Cyclotron On
  |B| ≃ 370 µT at strongest points
  MSR height above the floor: 1 m
  Z axis is height from floor

## Slide 8: Opera scalar potential applied to boundary/ result
Max field is 0.5 T
Need An external B field compensation system
Surface B field (T)
1.3E-6
0.32
0.25
0.20
0.15
0.10
0.05
Opera FEA (static mu=1M)
Using B=350 µT Vertical Field
Max field is 0.3 T
Opera FEA (static mu=1M)
Using B=350 µT Actual Field
Cyclotron on
2E-6
0.52
0.42
0.34
0.26
0.18
0.1
Surface B field (T)

## Slide 9: Summary
Going to measure the nEDM
Going to use an MSR at TRIUMF
Mapped the magnetic field
Need An external B field compensation system

## Slide 11
Backup slides

## Slide 12: Decomposition into spherical harmonics
Use Decomposition as found in C. Abel et.al, Phys. Rev. A 99, 042112 (2019)
Where Πi,l,m are the x, y, and z derivatives of the scalar basis functions. Gl,m are then the coefficients that describe the magnetic field in the volume by gradient.
Fit our magnetic Field maps up to 3rd order in B
In free space,

## Slide 13: Comparison of B field from GLMs to measured field maps
Python
Opera
Cyclotron off

## Slide 14: Magnetic Permeability