## Slide 1
Damon Todd
Yubin Hu, Alex Kireef, Wenhan Sun, Marco Salathe,
Janilee Benitez, Heather Crawford, Victor Watson, Larry Phair
Lawrence Berkeley National Laboratory
September 22, 2023
Using machine learning algorithms for ECR tuning and physics studies

## Slide 2: What we mean by machine learning (here, at least)
A data point: feature vector and a label (e.g. x=jpg, y=”cat”)
Data set: collection of feature vectors and labels
Goal:
Have computer learn from the data set so it can predict the label for a new vector
cat
cow
cat
cow
???

## Slide 3: Using machine learning with electron cyclotron resonance (ECR) ion sources
Possible elements:
Set parameters (magnetic fields, biased disk voltage, RF power, etc.
Measured quantities (biased disk current, x-ray flux, drain current, etc.)
A measured parameter (e.g., ion species current, emittance, etc.)
Machine learning:
Measure many vector/label pairs then predict a continuous variable label
Possible goals:
Maximize beam current
Minimize emittance
Maintain stability
Understand source physics???

## Slide 4: Our test bed: VENUS
VENUS
LBNL
VENUS:
World’s first fully-superconducting ECR ion source designed for 28 GHz operation
Injector for LBNL’s 88” Cyclotron
Prototype ECR for FRIB
RF: 28+18 GHz (up to 10+2 kW)
Maximum axial fields: 4 T
Example beams:
> 4.7 mA O6+, > 20 mA He+ from source
> 2 pμA, 5 MeV/u 48Ca11+ from Cyclotron for superheavy element research
197Au61+ extracted from cyclotron

## Slide 5: VENUS primary control and diagnostic parameters
ion beam
extraction coil
injection coil
middle coil (not shown)
sextupole
plasma
material:
gases, ovens, sputtering
RF:
18 and 28 GHz
biased disk
(not shown)
Diagnostics:
Faraday cup/analyzing magnet
Emittance scanner
Drain, bias currents
Cryostat x-ray load
Etc.

## Slide 6: VENUS operation and data collection
Operator
Controls
Diagnostics
Panelview + PLC
(Programmable Logic Controller)
VENUS
Computer control through PLC:
Advantage: Exploit 2 decades of safely logic in PLC
Disadvantage: Slow ~3 Hz communication

## Slide 7: What we have done
Collected over 1.5 years of all VENUS control and primary diagnostic data
Performed thousands of hours of computer-only control and optimization of VENUS without human interaction
This work is supported by the U.S. Department of Energy, Office of Science, Nuclear Physics program under Award Number DE-FOA-0002490.

## Slide 8: VENUS logging
Analyzing dipole field measurement dropped every ~13.5 hours for ~6 minutes with no beam effect
Incorrect calculated M/Q seen occasionally by cyclotron operators, but not discussed
Solution:
Incorrect ground for liquid nitrogen (LN) system affects hall probe during LN fills
Note:
This isn’t machine learning, but we are currently using the monitoring of VENUS for neural network training

## Slide 9: Computer-driven operation: baking
Blessing/curse of ECRIS: ion beams produced from any materials reaching plasma that don’t destroy the plasma
Adsorbed particles later desorbed:
Residues from previous work or handling of inner surfaces
Water, salts, etc. on all surfaces after atmosphere exposure
Prior beam materials (e.g., metals)
Plasma chamber desorption can be sped up using plasma as heating element
Acknowledgement...

## Slide 10: Computer-driven baking
Two-step process:
Source to full RF power
  Monitor pressure, currents, etc. and raise safely
  Time significantly reduced as the computer is persistent
“Slosh” plasma around to bake plasma chamber
  Change confining fields and add gas when desorption rate drops
  Every 6th change return to “base” fields and perform charge state distribution to monitor evolution

## Slide 11: Charge state distributions (CSDs) give feedback on progress
Carbon
Oxygen
Last ~6 bakes have been fully automatic
Not machine learning, but data will be used to produce more efficient baking technique (ML!)

## Slide 12: Problem: find extremum of an unknown function when measurements are expensive
Find maximum of unknown function:
Measurements are expensive (time, $, etc.)
Four points have been measured
Where to search next?
Method:
Fit curve between points
Find maximum
Repeat
Four curves
Ten curves
One hundred curves
Lots of curves

## Slide 13: Bayesian optimization basically follows a path like this
Notes:
Next search point has highest probability of improvement though other points may have curves with higher extrema
Path to maximum isn’t always smooth

## Slide 14: Full Bayesian Optimization of 124Xe37+
Parameter | Min | Max
Bias voltage [V] | 40 | 105
Oxygen valve | 11.6 | 12.5
Xenon valve | 8.0 | 13.0
Inj coil [A] | 185.6 | 186.0
Ext coil [A] | 136.6 | 136.8
Mid coil [A] | 152.0 | 152.3
Sext coil [A] | 430.3 | 430.5
18 GHz [kW] | 1.4 | 1.8
28 GHz [kW] | 5.2 | 6.0
(~5 min/point)
This is machine learning!
Records:
VENUS: ~40 μA
SECRAL II: ~50 μA
(fields and power much more conservative than LBL records)

## Slide 15: Augmenting search based on cost
Say cost of search increases with x
We could create a “cost function” where we only search a large x point if the likelihood of improvement is over some threshold
For superconducting sources:
Coils: Δtchange ➞ minutes
Bias: Δtchange ➞ seconds
We have employed cost functions to judiciously use coil changes

## Slide 16: Using cost function in VENUS optimizations
Proof of principle:
Three parameter optimization of O7+ (biased disk, middle solenoid, and oxygen valve) can be tens of minutes faster using cost function
Peak current found can be lower as it is a less-thorough search
We will continue to explore this optimization technique
O7+
Full optimization time

## Slide 17: Optimizing a little more like a human
Many human records are achieved by using a cost-function-like approach:
Coils are slow, so find a pretty good solution and work from there
Note:
When instabilities or too low pressures encountered, low current is recorded
Parameter | Min | Max
Bias voltage [V] | 25 | 65
Oxygen valve | 11.5 | 12.0
Xenon valve | 9.0 | 12.0
18 GHz [kW] | 1.20 | 1.80

## Slide 18: Are there trends we can see?
time [hours]
Code not being as smart as a person would be
Can change this by adjusting fitting functions

## Slide 19: Exploiting patience of computer
~5 minutes between optimization as computer performed charge state distribution (CSD) measurement at each step (~1-2 minutes)
Most human optimizations are missing CSD information
These four runs had almost identical 124Xe37+ currents but their CSDs were dramatically different
This information will be fed into neural network efforts for general plasma understanding

## Slide 20: Lots of data
Explore solenoid coil operation space using only 18 GHz
Difficult to discern any pattern or trend
plasma
beam
extraction coil
injection coil
middle coil (not shown)
inj. coil [A]
Inj. coil [A]

## Slide 21: Lots and lots of data: explore bias voltage and pressure influence
Two parameter space exploration:
Biased disk voltage
Oxygen gas flow
3695 configurations, averaged over 30 measurements each

## Slide 22: Summary and looking forward
We have thousands of hours of full-computer-control of VENUS for baking and beam optimization (limited by VENUS availability)
Our careful attempts at 124Xe37+ optimization are about 20% of record beams. Will push harder when backup chamber is complete
Team has begun on neural network to understand what leads to long-term instability
Next steps:
  Speed up data taking, CSDs, etc.
  Optimize optimization to transition away from exploring with time
  Use neural network to improve understanding and control long-term instabilities, improve baking, etc.
Thank you!!

## Slide 23
Backup slides

## Slide 24: Relations of maxima

## Slide 25: Visualization of unstable operation region
Two parameter space exploration:
Biased disk voltage
Oxygen gas flow
3695 configurations